## Slide 1: 3D Ultrasound for Scoliosis Quantification
Ben Church

## Slide 2: Contents
Background
  Scoliosis
  Assessment
  3D ultrasound
Ungi2014
Cheung2015

## Slide 3: Background - Scoliosis
Pathological spinal curvature
Develops during growth
Must be monitored

## Slide 4: Background – Assessment
X-ray imaging is the gold-standard
Measure Cobb angle from X-ray, quantifying severity

## Slide 5: Background – 3D ultrasound
Risks of cumulative X-ray exposure have motivated use of 3D ultrasound
Provides landmark locations rather than macroscopic visualization
Curvature extracted from landmark locations

## Slide 6: Ungi2014
Spinal Curvature Measurement by Tracked Ultrasound Snapshots

## Slide 7: Ungi2014 – Motivation
Health risks of repeated X-ray exposure
Cost and inaccessibility of MRI
Inaccuracy of palpation and surface topography

## Slide 8: Ungi2014 – Method
Scan phantom models with image in sagittal plane
Locate transverse processes
Extract curvature

## Slide 9: Ungi2014 – Validation
Examined correlation between their method and Cobb’s
Examined average errors and error ranges

## Slide 10: Ungi2014 – Results
R = 0.998 for adult model
R = 0.977 for pediatric model
Angle errors < 3.1° < 5°

## Slide 11: Ungi2014 – Critique
Phantom images unrealistically clear
Just used two phantom models
In vivo studies required

## Slide 12: Cheung2015
Freehand three-dimensional ultrasound system for assessment of scoliosis

## Slide 13: Cheung2015 – Motivation
Health risks of X-ray
Limitations of other methods described by [Ungi2014]
In vivo study required

## Slide 14: Cheung2015 – Method
Scan live patients with wide probe
Image in axial plane
Locate landmarks in software-generated model
Extract curvature using TPs, and TPs + SAPs

## Slide 15: Cheung2015 – Validation
Compared their method’s angle to Cobb angle
Examined intra and inter-user variability

## Slide 16: Cheung2015 – Results
X-ray Cobb angles ranged 1.9° to 29.9°
TP method ranged 0.48° to 20.7°
  R2 = 0.68
TP + SAP method ranged 0.8° to 25.9°
  R2 = 0.86

## Slide 17: Cheung2015 – Critique
Cobb angle range only 1.9° to 29.9°
  Mild to moderate
Unable to locate some landmarks
  Used alternative landmarks
  Discarded data
Method for dealing with imperfect data required

## Slide 18: Thank you!
Questions?

## Slide 19: References
[Ungi2014] T. Ungi, F. King, M. Kempston, Z. Keri, A. Lasso, P. Mousavi, J. Rudan, D. P. Borschneck, and G. Fichtinger, “Spinal Curvature Measurement by Tracked Ultrasound Snapshots”, Ultrasound in Medicine and Biology 2014; 40(2):447-454.
[Cheung2015] C.-W. J. Cheung, G.-Q. Zhou, S.-Y. Law, K.-L. Lai, W.-W. Jiang, and Y.-P. Zheng, “Freehand three- dimensional ultrasound system for assessment of scoliosis”, Journal of Orthopaedic Translation 2015; 3:123-133.